GNSS Data Processing

Technical & Verification Document

| Field | Detail |
| --- | --- |
| Product | GNSS Rover / Base Deformation Monitoring System |
| Prepared by | Encardio ________________ |
| Date | ________________ |
| Customer | ________________ |
| Site / Unit ID | ________________ |

# Contents

# 1. Purpose and Scope

This document describes, step by step and with the exact mathematics implemented in the firmware, how the system processes GNSS data — from the raw $GPGGA sentence in WGS-84 up to the deviation values reported by the Rover and the surveyed position reported by the Base.

It is written so that the customer can independently reproduce every value by hand (or in a spreadsheet) and compare it against the values reported by our software. Section 8 lists the constants, and Sections 2–6 each include a fully worked numeric example with exact intermediate results.

# 2. Reference Ellipsoid and Constants (WGS-84)

All conversions use the WGS-84 ellipsoid:

| Symbol | Meaning | Value (as used in code) |
| --- | --- | --- |
| a | Semi-major axis | 6378137.0 m |
| f | Flattening | 1 / 298.257223563 |
| e² | First eccentricity squared = f(2 − f) | 0.0066943799901413165 |

# 3. Step 1 — Reception of GGA Data (WGS-84)

The GNSS receiver outputs standard NMEA $GPGGA sentences once per second (1 Hz). Example:

$GPGGA,120000.00,2833.1234560,N,07712.5678900,E,4,18,0.8,213.500,M,...

Fields used by the software:

| Field # | Content | Example |
| --- | --- | --- |
| 2 | Latitude ddmm.mmmmmmm | 2833.1234560 |
| 3 | N/S hemisphere | N |
| 4 | Longitude dddmm.mmmmmmm | 07712.5678900 |
| 5 | E/W hemisphere | E |
| 6 | Fix quality | 4 (RTK fixed) |
| 7 | Satellites in use | 18 |
| 8 | HDOP | 0.8 |
| 9 | Altitude (above MSL) | 213.500 |

## 3.1 Data-quality acceptance gates

A fix is only accepted for processing if:

| Parameter | Rover requirement | Base survey requirement |
| --- | --- | --- |
| Fix quality (field 6) | = 4 (RTK Fixed only) | ≠ 0 (any valid fix) |
| Satellites (field 7) | ≥ 15 | ≥ 15 |
| HDOP (field 8) | ≤ 2.0 | ≤ 2.0 |
| Lat/Lon/Alt | numeric, non-NaN, not (0,0) | same |

Fixes failing any gate are discarded and never enter the average.

# 4. Step 2 — Conversion of GGA (Geodetic) to ECEF

## 4.1 NMEA degrees-minutes → decimal degrees

For a latitude/longitude field in ddmm.mmmm form:

deg     = floor(value / 100)

minutes = value - (deg * 100)

decimal = deg + minutes / 60

(if hemisphere is S or W, negate the result)

## 4.2 Geodetic (decimal degrees, ellipsoidal height) → ECEF (metres)

With latitude φ and longitude λ in radians, and height h:

N = a / sqrt(1 - e^2 * sin^2(φ))     (prime-vertical radius)

X = (N + h) * cos(φ) * cos(λ)

Y = (N + h) * cos(φ) * sin(λ)

Z = (N * (1 - e^2) + h) * sin(φ)

Note on height (h): the software uses GGA field 9 (altitude) directly as h. The geoid–ellipsoid separation is not added. Because the Rover reports relative deviation (Section 6), this constant offset cancels out and does not affect the reported deviation. For manual verification, use GGA field 9 directly, exactly as the software does.

## 4.3 Worked Example A — one fix converted to ECEF

Input: latitude = 2833.1234560 N, longitude = 07712.5678900 E, altitude = 213.500 m

Step 1 — decimal degrees:

lat = 28 + 33.1234560/60 = 28.5520576000°

lon = 77 + 12.5678900/60 = 77.2094648333°

Step 2 — ECEF (using the constants in Section 2):

N = 6383019.591553 m

X = 1241302.544771 m

Y = 5467794.230140 m

Z = 3030487.891673 m

(These are the exact values the software computes for this input.)

# 5. Step 3 — Base Station Processing (Position Survey)

The Base establishes its own fixed reference position by averaging many accepted fixes, with robust outlier rejection.

## 5.1 Two-level averaging structure

- Each accepted fix is converted to ECEF (X, Y, Z) and added to a window buffer.

- Every 600 samples (≈10 minutes at 1 Hz) the buffer is reduced to one window-average using the MAD filter (5.2–5.3).

- Windows required for a complete survey = base_reading_points / 600  (e.g. 6 hr → 21600, so 21600/600 = 36 windows).

- When enough windows exist, all windows are combined into the final Base position by a weighted average (5.4), then converted back to latitude/longitude/height (5.5).

## 5.2 MAD (Median Absolute Deviation) outlier filter

For a set of values (applied independently to each of X, Y, Z, but the keep/reject decision is joint):

med   = median(values)

MAD   = median( |value_i - med| )

sigma = 1.4826 * MAD   (robust estimate of standard deviation)

A sample i is kept only if it is within the cutoff on all three axes simultaneously:

|X_i - medX| <= k*sigmaX  AND

|Y_i - medY| <= k*sigmaY  AND

|Z_i - medZ| <= k*sigmaZ

(cutoff per axis = max(1e-6, k*sigma))

## 5.3 Adaptive k (retain ≥ 70 %)

The multiplier k starts at 1.0 and increases in steps of 0.2 (max 10.0) until at least 70 % of samples are retained:

k = 1.0

while (fraction_kept < 0.70 and k <= 10.0):

k = k + 0.2

## 5.4 Window average and weighted combination

- Within a window: kept samples are averaged (simple mean) per axis; the count of kept samples is stored with the window.

- Across windows: the final position is a count-weighted mean: X_final = Σ(meanX_i × count_i) / Σ(count_i)  (same for Y, Z).

## 5.5 ECEF → Geodetic (final Base coordinate)

The final averaged ECEF is converted back to latitude/longitude/height using the standard iterative method:

λ  = atan2(Y, X)

p  = sqrt(X^2 + Y^2)

φ0 = atan2(Z, p*(1 - e^2))                 (initial estimate)

repeat until |φ - φ_prev| < 1e-12 :

N = a / sqrt(1 - e^2 * sin^2(φ))

h = p / cos(φ) - N

φ = atan2( Z, p*(1 - e^2 * N/(N + h)) )

h = p / cos(φ) - N                          (final height)

## 5.6 Worked Example B — MAD filter on one window (X-axis shown)

Seven X-values (metres), six clustered and one gross outlier:

1241226.100, 1241226.140, 1241226.090, 1241226.160,

1241226.120, 1241226.130, 1241228.900   <- outlier

median (X)         = 1241226.130000

MAD                = 0.030000

sigma = 1.4826*MAD = 0.044478

k = 1.0 -> cutoff = 0.044478 -> kept = 6/7 = 85.7%  >= 70%  -> STOP

kept values     = 1241226.100 ... 1241226.130  (outlier rejected)

window mean (X) = 1241226.123333   count = 6

(The same procedure runs simultaneously on Y and Z; a sample survives only if it passes on all three axes.)

## 5.7 Worked Example C — weighted average across windows (X-axis)

| Window | mean X (m) | count |
| --- | --- | --- |
| 1 | 1241226.123 | 98 |
| 2 | 1241226.118 | 100 |
| 3 | 1241226.130 | 95 |

X_final = (1241226.123*98 + 1241226.118*100 + 1241226.130*95) / (98+100+95)

= 363679254.2040 / 293

= 1241226.123563 m

# 6. Step 4 — Rover Processing (Deviation Reporting)

The Rover measures how much its position has moved relative to its own initial reference reading.

## 6.1 Accepted data

Only RTK-Fixed fixes (quality = 4) meeting the gates in Section 3.1 are used. Each is converted to ECEF (Section 4).

## 6.2 Moving-average structure

- Each accepted ECEF fix is added to a window buffer.

- Every 100 samples the buffer is reduced to one window-average using the same MAD filter (5.2–5.4).

- The moving average uses the most recent N windows.

Continuous mode:  N = base_line_duration / 100

Burst mode:       N = (sessions * burst_window_seconds) / 100

sessions = baseline_seconds / burst_interval_seconds

Example (Burst, 6 hr baseline, 2 hr interval, 15 min window):

sessions = 6/2 = 3;  N = (3 * 900)/100 = 27 windows (~ last 2700 fixes)

## 6.3 Initial reading (baseline)

The baseline is the Rover's reference 'zero' position. It is computed once — the first time the window buffer reaches N windows — using the same weighted average, then stored and never changed (until a manual reset). All subsequent deviations are measured against it.

## 6.4 Deviation calculation and reporting

For each reporting interval, the current moving-average ECEF is differenced against the stored baseline ECEF:

dX = X_current_average - X_baseline

dY = Y_current_average - Y_baseline

dZ = Z_current_average - Z_baseline

The result is then scaled to the reporting unit:

| Unit | Factor | Decimals |
| --- | --- | --- |
| mm | × 1000 | 1 |
| cm | × 100 | 2 |
| inch | × 39.3701 | 2 |
| feet | × 3.28084 | 4 |

IMPORTANT — coordinate frame of the deviation columns: the reported columns dNORTHING, dEASTING, dALTITUDE are the ECEF component differences dX, dY, dZ respectively (the software carries ECEF X/Y/Z through these internal names). They are NOT local North/East/Up. To reproduce them, difference the ECEF coordinates as shown above. The total displacement magnitude sqrt(dX²+dY²+dZ²) is frame-independent and can also be checked directly.

## 6.5 Worked Example D — deviation reported in mm

| Component | Baseline (m) | Current avg (m) | Difference (m) | Reported (mm) |
| --- | --- | --- | --- | --- |
| dX (dNORTHING) | 1241226.120 | 1241226.125 | +0.005000 | +5.0 |
| dY (dEASTING) | 5467447.300 | 5467447.318 | +0.018000 | +18.0 |
| dZ (dALTITUDE) | 3030380.500 | 3030380.492 | -0.008000 | -8.0 |

Total 3-D displacement = sqrt(5² + 18² + 8²) = 20.2 mm.

## 6.6 CSV output columns

DATE/TIME, dNORTHING(unit), dEASTING(unit), dALTITUDE(unit),

NORTHING, EASTING, ALTITUDE(m), SATELLITES, BATTERY(V), TEMPERATURE(DEG)

- dNORTHING/dEASTING/dALTITUDE = ECEF dX/dY/dZ (Section 6.4), in the selected unit.

- NORTHING/EASTING/ALTITUDE(m) = current averaged absolute position, converted back to lat/lon/height via Section 5.5.

- SATELLITES = satellites used in the last accepted fix.

# 7. Summary of the End-to-End Pipeline

$GPGGA (WGS-84)                                   [Section 3]

|  quality gates: fix, sats >= 15, HDOP <= 2.0

v

Decimal degrees -> ECEF (X,Y,Z)                   [Section 4]

|

+-- BASE:  600-sample windows -> MAD (>=70%) [Section 5]

|          -> count-weighted mean of windows

|          -> ECEF->geodetic = fixed Base position

|

+-- ROVER: 100-sample windows -> MAD (>=70%) [Section 6]

-> count-weighted moving average (N windows)

-> baseline captured once (initial reading)

-> deviation = current - baseline (ECEF dX,dY,dZ)

-> x unit factor -> reported (mm/cm/inch/feet)

# 8. Appendix — Constants and Verification Checklist

## Constants

a  = 6378137.0

f  = 1 / 298.257223563

e^2 = 0.0066943799901413165

MAD scale factor = 1.4826

MAD retain target = 70%, k from 1.0 step 0.2 (max 10.0)

Base window size  = 600 samples

Rover window size = 100 samples

## Manual verification checklist

- Convert a chosen $GPGGA to decimal degrees (§4.1) and to ECEF (§4.2). Compare with Example A.

- Take a set of accepted fixes; compute median, MAD, sigma, apply the ≥70% keep rule (§5.2–5.3, Example B).

- Compute the count-weighted window average (§5.4, Example C).

- For the Rover, difference the current average against the stored baseline in ECEF and apply the unit factor (§6.4, Example D).

- Cross-check the total displacement sqrt(dX²+dY²+dZ²) — independent of coordinate frame.

All formulas in this document correspond exactly to the firmware implementation. For any value that does not match, please share the raw $GPGGA records and the reported CSV row, and we will reconcile it line by line.